"These data are property of Rosetta Inpharmatics, Inc., 12040 - 115th Ave. NE, Kirkland, WA 98034-6900.  www.rii.com  All rights reserved.  Your use of this data may only occur pursuant to a written license agreement between you (or your employer) and Rosetta Inpharmatics."

Description of methods

The cDNA microarray experimental procedures (arraying, hybridization, and spot quantitation) in this study followed those outlined in Schena et al. (1995), and more precisely followed methods described in Marton et al. (1998).  A variety of aspects of manufacturing and using cDNA microarrays are described in detail at:

http://cmgm.stanford.edu/pbrown/

Array manufacture

A custom-built spotter with 16 quill pens was used to deposit PCR products corresponding to >6000 yeast genes within a 2 x 2 cm2 area on standard laboratory glass slides treated with poly-lysine.  Spotted slides were post-processed using standard protocols (heat -> UV -> block -> boil -> ethanol rinse -> dry).

Sample preparation

Growth conditions are described in the text and in "Experimental Procedures".  RNA was extracted from yeast cultures by hot phenol and vortexing.  mRNA was purified by two passes over oligo-dT cellulose (NEB).  Cy3 and Cy5 were incorporated into cDNA as described (Marton et al., 1998).

Hybridization

Samples were hybridized to microarrays under an ordinary 22 mm2 glass coverslip for 12-16 hours at 62.5 degrees C in 3x SSC with 0.75 mg/ml poly-A DNA and 0.2% SDS; the total hybridization volume was ~20 l.  Following hybridization, slides were washed in SSC/SDS solutions as described (Marton et al., 1998).

Scanning

Hybridized arrays were scanned on one of two scanning platforms that yielded nearly identical results:  a General Scanning ScanArray 3000 (see www.genscan.com), which is a confocal laser scanner, or an API Arrayworx (see www.api.com), which is a CCD device.  Scanning parameters (light intensity and detection levels) were set such that the brightest spots were just below saturation.

Image analysis and quantitation

Image analysis is described in Marton et al., 1998.  First, a grid is superimposed over the scan images.  The center of each spot is determined automatically and the correct grid location is confirmed by the user.  Prior to quantitation, the user edits the grid to exclude spots with slide artifacts such as dust and salt residue.  Quantitation of ratios and intensities is carried out in three automated steps:  First, the background for each channel of each slide is determined; second, the intensity of each spot over background is measured in each channel (see Figure 1), and third, the two channels are normalized to total signal intensity, i.e. rather than measuring absolute transcript abundances, or using one or a few 'housekeeping' genes to normalize, ratios are determined simply assuming that the total transcript abundance is unchanged.  This approach compensates for possible biases in intensity resulting from inefficiencies in sample extraction and labeling, and is ideal for pattern matching among profiles.  Spots saturated in either channel are flagged.  Image processing also involves detrending artifactual ratio patterns correlating with parameters such as transcript length and position on the array, to compensate for differing properties of Cy3- and Cy5-labeled transcripts, and slight variations in scanner focus.  Following quantitation, data tables including error estimates and P-values are generated using the formulas presented in the 'Error model' document.

Figure 1:  Spot quantitation.  The pixels within the white circle are quantitated.  Outlier pixels within the white circle will not be included in the quantitation.

Reproducibility of profiles

To confirm the reproducibility of expression profiles from mutants, 152 mutants were profiled in duplicate (i.e., cells starting from the same glycerol stock were taken through the growth and analysis procedure from start to finish).  The Experiment list indicates which mutants were repeated; for repeated mutants, each data points is a composite of all four slides from the same mutant (since each individual comparison was hybridized to two identical arrays in 'fluor reversal').  The correlation between duplicates is plotted in Figure 2 as a function of the number of significantly regulated genes.  We conclude from these data that overall reproducibility is qute high, considering that a correlation of 0.4 among 5 or more genes (using the gene-specific error model) is generally sufficient to drive association in clustering analysis.

Figure 2:  Reproducibility.   The first plot employs a simple 2-fold cutoff as the significance measure; the second uses P < 0.01 with measurement error only; the third uses P < 0.01 with the gene-specific error model.  Experiments with no significant genes are excluded.